Syringe Pump
Specification:-

1. INFUSION FLOW RATE RANGE 0.1-1200ml/h (depending on the syringe capacity).

2. FLOW RATE ACCURACY  +1% on mechanism, + 2% on Syringes.

3. SYRINGE CAPACITIES 5, 10, 20, 30/35, 50/60 CC.

4. VOLUME/DOSE INFUSEDVolume: 0, 1-999,9ml.Dose: starts at 0.1 ml.

5. INFUSION MODES ml/h.
· Dose rate.

Dose units;ng,ug,mg,cal,kcal,U,kU,mmol,mol,Kmol/kg.

Time units: 24h/h/min.

Volume or dose/time:0.1-99.9 units: 00h01-96h00.

Volume limit: 0.1-999.9 ml.

6. INFUSED Volume: 0, 1-999,9ml:Dose:starts at 0.1ml.
7. BOLUS Manual bolus: Rate: 50 ml/h 50-1200ml/h: 50ml/h.

8. Programmed bolus (dose or volume/time): 0.1-99.9 units/0.1-1200 ml time auto –calculation. 
Drug labeling: infusion with the drug names. PAUSE Programmable from 1 minute to 24 hours, increments from minute to minute.
9. DRUG LABELLING Display of the drug name based on a list of 50 pre-set drug names.

PRESSURE MANAGEMENT PRESSURE MODES 2 modes available: variable or 3 pre-set levels-from 100 to 900 mm. Hg increment.

DPS the Dynamic Pressure System-DPS-warns of pressure variations. A risk of obstruction or a possible leak in the infusion line can thus be anticipated.

10. PRESSURE MONITORING Graphic representation of the pressure limit.

11. ALARMS/PRE-ALARMS/SECURITYTECHNICAL SPECIFICATIONS MANUAL PUSHER Protection for the ongoing infusion.

12. ALARMS/PRE-ALARMS/SECURITY PUMP STATUS Green for infusion in progress, orange for pre-alarm and warning, red for alarm.

13. Visibility up to 5 meters. All alarms are expressed by means of light indicators, written words, pictograms and sound beep.
14. INFUSION CONTROL Occlusion pressure pre-alarm, occlusion pressure alarm, end of(from 1 to 30 minutes) before the end of infusion, volume limit pre-alarm, volume limit alarm, keypad locking, hard and soft flow rate limits. low battery pre- alarm ,discharges battery alarm, battery capacity display in hours and minutes, unconfirmed programming, technical malfunction alarm(auto-test, rotation), drive system advance check, Watchdog check, communication connection failure.
15. ALARMS/PRE-ALARMS/SECURITY TECHNICAL SPECIFICATION MANUAL PUSHER Protection for the ongoing infusion.”Push-Guard”. SCREEN Blue graphic LCD monochrome, size 66 mm x 33 mm (128x 64 pixels.)

16. Battery back up-Minimum 4 Hrs

17. Accessories (1) Power Cord

                          (2).Pole clamp.

